GOVERNMENT OF INDIA
MINISTRY OF CONSUMER AFFAIRS, FOOD & PUBLIC DISTRIBUTION
DEPARTMENT OF CONSUMER AFFAIRS — LEGAL METROLOGY DIVISION

TYPE EVALUATION TEST REPORT
Non-Automatic Weighing Instrument as per OIML R 76-1:2006 (E)

	Report Number: REP-2026-OIML-002-FAIL
	Date: 08-Oct-2026

	Session Number: TS-2026-002-FAIL-DEMO
	

	Overall Result: FAIL
	Standard: OIML R 76-1:2006 (E)



1. Applicant & Manufacturer Information
	Applicant Name
	Demo WeighTech Pvt. Ltd.

	Applicant Address
	Plot 45, Okhla Industrial Area Phase III, New Delhi 110020, India

	Manufacturer Name
	Demo WeighTech Pvt. Ltd.

	Manufacturer Address
	Plot 45, Okhla Industrial Area Phase III, New Delhi 110020, India



2. Technical & Metrological Characteristics
	Model & Serial Number
	MW-3000 / S/N: MW3000-2026-IND-001

	Accuracy Class
	Class III

	Max & Min Capacity
	Max = 3000.0 kg, Min = 20.0 kg

	Scale Intervals (e, d, n)
	e = 1.0 kg, d = 1.0 kg, n = 3000

	Load Receptor / Supports
	Platform (4 support points)

	Power & Temp Range
	230 V AC, 50 Hz | -10.0°C to 40.0°C



3. OIML R-76 Test Applicability & Coverage Matrix
	Code
	Clause
	Test Procedure
	Applicability
	Execution
	Status

	A.1
	Clause A.1
	Administrative Examination
	APPLICABLE
	NOT TESTED
	PENDING

	A.2
	Clause A.2
	Compare Construction with Documentation
	APPLICABLE
	NOT TESTED
	PENDING

	A.3
	Clause A.3
	Initial Examination
	APPLICABLE
	NOT TESTED
	PENDING

	A.4.2.1
	Clause 4.5.1 / A.4.2.1
	Range of Zero-Setting
	APPLICABLE
	NOT TESTED
	PENDING

	A.4.2.3
	Clause A.4.2.3
	Accuracy of Zero-Setting
	APPLICABLE
	EXECUTED
	PASS

	A.4.4.1
	Clause A.4.4.1 & A.4.4.3
	Weighing Performance Test
	APPLICABLE
	EXECUTED
	PASS

	A.4.7
	Clause A.4.7
	Eccentricity Test
	APPLICABLE
	EXECUTED
	FAIL

	A.4.8
	Clause A.4.8
	Discrimination Test
	APPLICABLE
	NOT TESTED
	PENDING

	A.4.10
	Clause A.4.10
	Repeatability Test
	APPLICABLE
	NOT TESTED
	PENDING

	A.4.11.1
	Clause A.4.11.1
	Creep Test
	APPLICABLE
	NOT TESTED
	PENDING

	A.4.11.2
	Clause A.4.11.2
	Zero Return Test
	APPLICABLE
	NOT TESTED
	PENDING

	A.5.1
	Clause A.5.1
	Tilting Test
	APPLICABLE
	NOT TESTED
	PENDING

	A.5.3.1
	Clause A.5.3.1
	Static Temperatures Test
	APPLICABLE
	NOT TESTED
	PENDING

	A.5.3.2
	Clause A.5.3.2
	Temperature Effect on No-Load Indication
	APPLICABLE
	NOT TESTED
	PENDING

	A.5.4
	Clause A.5.4
	Voltage Variations Test
	APPLICABLE
	NOT TESTED
	PENDING

	A.6
	Clause A.6
	Endurance Test
	APPLICABLE
	NOT TESTED
	PENDING

	B.2
	Clause B.2
	Damp Heat, Steady State
	APPLICABLE
	NOT TESTED
	PENDING

	B.3
	Clause B.3
	Performance Tests for Disturbances (EMC / ESD / Dips)
	APPLICABLE
	NOT TESTED
	PENDING

	B.4
	Clause B.4
	Span Stability Test
	APPLICABLE
	NOT TESTED
	PENDING

	Annex G
	Annex G (Cl. G.1 - G.4)
	Software-Controlled Instrument Examination
	APPLICABLE
	NOT TESTED
	PENDING



4. Summary of Executed Test Results & Metrological Compliance
	Clause
	Test Description
	Observed Result
	Permissible Limit
	Result

	Clause A.4.2.3
	Accuracy of Zero-Setting
	E0 = +0.000 kg
	Limit = ±0.250 kg (0.25e)
	PASS

	Clause A.4.4.1 & A.4.4.3
	Weighing Performance Test
	1 points evaluated (Max |Ec| = 0.000 kg)
	Tiered MPE as per Table 6 for Class III
	PASS

	Clause A.4.7
	Eccentricity Test
	Max eccentricity error = 2.400 kg
	Limit = ±1.000 kg
	FAIL




5. Review & Approval Signatures

	Evaluated by:

___________________
Test Engineer
Date: 08-Oct-2026
	Reviewed by:

___________________
Technical Reviewer
Status: UNDER REVIEW (Pending Review)
	Approved by:

___________________
Director / Lab In-Charge
Status: Pending Final Approval



DISCLAIMER: This prototype test report is generated by METRIREPORT for SIH 2026. It is not an official model approval certificate under the Legal Metrology Act.




